| Short Title | VO2 Mouthpiece Cleaning |
| --- | --- |
| Effective Date | 3 February 2010 |
| Approved by REB | 3 February 2010 |
| Version Number | 1 |

* Always wear a lab coat and gloves when cleaning mouthpiece.  Use forceps to take mouthpiece out of liquid

1) Rinse mouthpiece with water, separate used mouthpiece from mouthpiece apparatus and place in Klenzyme Container #1, shake and let it sit for minimum 5 min (can be left longer if no chance to clean).

2) Rinse mouthpiece with water (Container #2) for 1 min.

3) Place mouthpiece into Cidex (Container #3) for 8 - 10 min.  **Leaving the mouthpiece soaking in Cidex will result in a build-up which is harmful for use, please carry a timer with you to remind you to remove from solution.  No longer than 10 min.

* Note: Cidex should be tested with a test strip prior to use.  If the strip fails the Cidex should be replaced before the next use. Clean container with Klenzyme prior to refilling with Cidex.  Please see manufacturers’ guidelines for correct dilutions/concentrations.  Expiration dates should be confirmed on container prior to use. The container should be dated when the Cidex has been changed.  Cidex is a harmful agent and the container must be sealed when not in use.

4) Rinse, rinse, rinse!!! mouthpiece thoroughly with fresh tap water.  Place mouthpiece in water container, shake container and let sit for 1 min. Empty water into sink and refill water at least 2 more times.

5) Dry mouthpiece by hand with sterile towel and place in the clean dry container (#4) for next use.

6) All clean mouthpieces should be stored in the sealed container # 4, do not leave uncovered.  This container must be stored carefully so that no re-contamination can occur.

*** Only assemble clean mouthpiece and breathing apparatus with clean or gloved hands***

Please see PI if you have any questions re: procedure, and ordering new supplies (well in advance of running out!).